## Slide 1: Handpiece Cleaner/Lubricator Prototyping at M3 Design
Siddharth Kurwa
skurwa.github.io

## Slide 2: Agenda
About Me
Project Background
Design Process
  Hardware Selection
  Low-Resolution Prototype
  Functional Prototype
  Looks-Like/Works-Like Prototype
Outcome
Questions

## Slide 3: A Few of My Favorite Things
Mediterranean Food
Zion National Park
Tinkering
Project Background
Hardware Selection
Low-Res Prototype
Functional Prototype
Looks-Like/Works-Like
About Me
Outcome

## Slide 4: Project Background
Client Need
  Field support for failed medical handpieces is expensive
  Handpiece cycle life reduced by poor manual maintenance
Problem Statement
  Design and build a cleaning/lubricating system to simplify handpiece maintenance
Timeline
  1.5 months
Team Roles
  2 ME – fluid system design, CAD, mechanical prototype
  1 ID - user workflow, product aesthetic
  1 EE/ME - verification, support, debug
  1 EE - hardware selection, software development, system testing
Client’s surgical handpiece products.
Project Background
Hardware Selection
Low-Res Prototype
Functional Prototype
Looks-Like/Works-Like
About Me
Outcome

## Slide 5: Product Externals
Cleaner and lubricant cartridges
Handpieces
Cartridge tray
Door
OLED display
Pushbuttons
Project Background
Hardware Selection
Low-Res Prototype
Functional Prototype
Looks-Like/Works-Like
About Me
Outcome

## Slide 6: My Scope
Select, design, and build the electronics and control system to meet mechanical and industrial design requirements in a 1.5 month timeline.
Project Background
Hardware Selection
Low-Res Prototype
Functional Prototype
Looks-Like/Works-Like
About Me
Outcome

## Slide 7: Requirements
User workflow map.
Source | Type | Requirement
Industrial Design | Sensing | Detect open/closed door
 | Sensing | Detect open/closed drawer
 | UI | OLED character display
 | UI | 2 pushbuttons
 | UX | 10 user workflow maps
Mechanical | Sensing | Detect tank overflow
 | Controls | Satisfy pump pressure/flow operating conditions
 | Volume | Fit within 12” x 4” x 6” volume
Table of design requirements.
Project Background
Hardware Selection
Low-Res Prototype
Functional Prototype
Looks-Like/Works-Like
About Me
Outcome

## Slide 8: Process
Hardware Selection
Low-Res Prototype
Functional Prototype
Looks-Like, Works-Like Prototype
Identify possible hardware options.
Bench test components to verify viability.
Build and test client deliverable.
Integrate subsystems to verify interaction.
Project Background
Hardware Selection
Low-Res Prototype
Functional Prototype
Looks-Like/Works-Like
About Me
Outcome

## Slide 9
Hardware Selection
Project Background
Hardware Selection
Low-Res Prototype
Functional Prototype
Looks-Like/Works-Like
About Me
Outcome

## Slide 10: Sensor Survey
Open/closed door and drawer
  Considered detent and reed switches – quickly chose reed switches for smaller package size
Liquid level switch
  Researched sensing principles (ultrasonic, capacitive, optical, pressure, magnetic) and found several options – narrowed down to below list
Working Principle | Source | Lead Time | Pros | Cons
Capacitive | DFRobot | 1-2 weeks | Cost; non-contact; package size | Limited resources; lead time
Capacitive | Carlo Gavazzi | 1 day | Non-contact; lead time; support | Cost; package size
Optical | Sparkfun | 1 day | Cost; lead time; package size | Contact
Comparison of liquid level switch solutions.
Project Background
Hardware Selection
Low-Res Prototype
Functional Prototype
Looks-Like/Works-Like
About Me
Outcome

## Slide 11: Pump Control Solution
Pump operating condition requirements
  Open handpieces: 30 GPH, <4 psi
  Restrictive handpieces: 0 GPH, <9 psi
Tested pump characteristics without control
  Open flow: 30 GPH, 3.6 psi
  Choked flow: 0 GPH, 16 psi
Made physical system model to identify control need
  Motor driver to satisfy requirement
Sense resistor controls current limiting
DRV8871 motor driver.
Output pressure head
K2
Motor torque
K1
Input current
Controller saturates input current
Open-loop control model of pump.
Project Background
Hardware Selection
Low-Res Prototype
Functional Prototype
Looks-Like/Works-Like
About Me
Outcome

## Slide 12
Hardware Selection
Low-Res Prototype
Project Background
Hardware Selection
Low-Res Prototype
Functional Prototype
Looks-Like/Works-Like
About Me
Outcome

## Slide 13: Sensor Function Checking
0.1” repeatability range
Carlo Gavazzi sensor
Level sensor bench test setup.
¼” shim
No shim
0.1” repeatability acceptable with margin based on tank CAD and projected sensor position
Test results.
Project Background
Hardware Selection
Low-Res Prototype
Functional Prototype
Looks-Like/Works-Like
About Me
Outcome

## Slide 14: Pump Control Testing
Iteratively determined current sense resistor to meet pressure requirements
Provided empirical model:
Motor driver
Pressure transducer
Flowmeter
Choke valve
Pump
Pump measurement setup.
Project Background
Hardware Selection
Low-Res Prototype
Functional Prototype
Looks-Like/Works-Like
About Me
Outcome
Motor driver outputs.

## Slide 15: State Machine Software
Arduino microcontroller for digital IO
Used simple components to simulate physical system
Initial prototype to drive LED output.
Incrementally added components.
Toggle switches
(sensors)
Relay module
OLED
display
LEDs
LCD display
Alarm LED output
Finished LED output
Run LED output
Idle LED output
State selector
Elapsed time
Previous state
Pushbutton interrupts
Toggle switch interrupts
Foundational state machine architecture.
State
Project Background
Hardware Selection
Low-Res Prototype
Functional Prototype
Looks-Like/Works-Like
About Me
Outcome

## Slide 16
Hardware Selection
Low-Res Prototype
Functional Prototype
Project Background
Hardware Selection
Low-Res Prototype
Functional Prototype
Looks-Like/Works-Like
About Me
Outcome

## Slide 17: State Machine Expansion
Hierarchical state machine to control sequencing of run-time processes
Handled nominal and edge cases with 11 states and event/interrupt hierarchy
Ex. edge case: enter run, trip door alarm, trip drawer alarm
Drawer alarm output
Door alarm output
Idle output
State selector, process selector
Elapsed time
Previous state
Pushbutton interrupts
Sensor interrupts
…
Run
Add cleaner
Close drain
Open drain
Pump fluid
…
Robust state machine architecture.
State
Process
Project Background
Hardware Selection
Low-Res Prototype
Functional Prototype
Looks-Like/Works-Like
About Me
Outcome

## Slide 18: System Integration
Validated pump performance in actual system
Supported ME in tuning process function timing and sequence
Functional prototype.
Project Background
Hardware Selection
Low-Res Prototype
Functional Prototype
Looks-Like/Works-Like
About Me
Outcome
Run agitator
Add cleaner
Fill tank
Determined process sequence.
Pump fluid
Drain tank
Run agitator
Add lubricant
Fill tank
Pump fluid
Drain tank
Drain tank

## Slide 19: System Debug
Issue
  Solenoid valves overheating – melting risk to acrylic manifold/reservoir
Context
  Solenoid directly connected to power supply, bang-bang control
Root Cause Analysis
  Found solenoid valve power rating on spec sheet (7W)
  Estimated power dissipation (16W) based on winding resistance
Solutions
  Limit solenoid valve current draw
    Add in-series resistance – cheaper but slower
    Swap relay to motor driver with current limiting feature – immediate but more expensive
Solenoid valve
Acrylic manifold
Solenoid valve and acrylic manifold interface.
Project Background
Hardware Selection
Low-Res Prototype
Functional Prototype
Looks-Like/Works-Like
About Me
Outcome

## Slide 20
Hardware Selection
Low-Res Prototype
Functional Prototype
Looks-Like, Works-Like Prototype
Project Background
Hardware Selection
Low-Res Prototype
Functional Prototype
Looks-Like/Works-Like
About Me
Outcome

## Slide 21: Looks-Like, Works-Like Prototype
Incorporated sensors, UI features, and reliable connections
User interface
Reed switch
Molex connectors
Adding final electronics features to system.
Powering on system with finalized components.
Project Background
Hardware Selection
Low-Res Prototype
Functional Prototype
Looks-Like/Works-Like
About Me
Outcome

## Slide 22: System Debug
Issue
  System freezes when pumps turn on
Context
  3 sensors and 2 pushbuttons use interrupt service routines (ISRs) with simple delay debounce routine
    If interrupt: wait 25 ms, then check state: ISR if high, exit if low
Root Cause Analysis
  Removed sensor response from logic – timing worked reliably
  Added interrupts back incrementally – timing froze again with both reed switches
  Possible causes: weak magnet, EMI signal noise, mechanical vibration
Solutions
  Software Patches
    Counter-based debounce routine to minimize time spent in ISR
    Change reed switches to polling architecture rather than event-based
  Hardware
    Stronger magnets
    Twisted-pair cabling of high amperage lines to minimize EMI
Reed switch.
Project Background
Hardware Selection
Low-Res Prototype
Functional Prototype
Looks-Like/Works-Like
About Me
Outcome

## Slide 23: Outcome
Project
  $200,000 billable project delivered on schedule
  Client demoing prototype to customers since mid-Sep without issue
  Client returned with 2 new projects
Personal
  Owned critical components and deadlines of multidisciplinary project (ME/EE/ID)
  Built tuneable controls architecture to handle every case
  Systematically debugged issues to identify root cause and corresponding solution
Project Background
Hardware Selection
Low-Res Prototype
Functional Prototype
Looks-Like/Works-Like
About Me
Outcome

## Slide 24: Questions?